# 6 Sterilization and Disinfection in the Laboratory

It is important to distinguish between sterilization and disinfection. Whereas sterilization results in destruction of all forms of microbial life, disinfection results in destruction of specific pathogenic microorganisms. A more detailed description of disinfection levels can be found in the Glossary at the back of this manual (Section 13).

## 6.1 Microbial Resistance to Physical and Chemical Agents

Microorganisms vary in their resistance to destruction by physical or chemical means. A disinfectant that destroys bacteria may be ineffective against viruses or fungi. There are differences in susceptibility between gram-negative and gram-positive bacteria, and sometimes even between strains of the same species. Bacterial spores are more resistant than vegetative forms, and non-enveloped, non-lipid-containing viruses respond differently than do viruses which have a lipid coating. Information on the susceptibilty of a particular microorganism to disinfectants and physical inactivation procedures can be found in the material safety data sheet (MSDS) for that agent. MSDSs provide additional details such as health hazards associated with the microorganism, mode of transmission, containment requirements and spill response procedures. The Environmental Safety Office has available, and can provide to individuals, MSDSs on a number of infectious microorganisms.

## 6.2 Physical Sterilants and Disinfectants

### 6.2.1 Heat Sterilization and Decontamination

Generally, sterilization is best achieved by physical methods such as steam or dry heat, which are less time-consuming and more reliable than chemical germicides. A summary of physical agents which employ heat for control of microorganisms can be found in Table 2. Of these physical procedures, steam autoclaving is the most practical option for the majority of laboratories for both sterilization and decontamination purposes.

Details on the use of an autoclave are given in Section 8.8.

| TABLE 2 - Outline of the properties of heat decontamination methods. For everyday laboratory purposes, autoclaving is the preferred method, unless the item cannot withstand the heat and/or moisture of autoclaving. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Principle/Conditions | Advantages | Disadvantages | Uses |
| Dry Heat | Thermal inactivation: destroys by oxidation | Non-corrosive Simple design and principle | Less effective than moist heat; requires longer times and/or higher temperatures | Materials that are damaged by, or are impenetrable to, moist heat |
| Hot Air Oven |  160-180?C for 2-4 hours |  penetrates water-insoluble materials (e.g., grease and oil)  less corrosive to metals and sharp instruments than steam |  slow diffusion, penetration  loading, packing critical to performance  not suitable for reusable plastics |  anhydrous materials, such as oils, greases and powders  laboratory glassware, instruments  closed containers |
| Red-heat Flame |  oxidation to ashes (burning) |  rapid |  initial contact with flame can produce a viable aerosol  possibility of accidental fire |  inoculating loops, needles |
| Incineration |  oxidation to ashes (burning)  1-60 minutes: temperatures may exceed 1000?C |  reduces volume of waste by up to 95% |  improper use may lead to emission of pathogens in smoke  requires transport of infectious waste  excess plastic (>20%) content reduces combustibility |  for decontamination of waste items prior to disposal in landfill |
| Moist Heat | Irreversible coagulation of (microbial) proteins | More rapid and more effective than dry heat |  |  |
| Pasteurization |  heating to below boiling point (generally 77?C) for up to 30 minutes |  can be used on heat sensitive liquids and medical devices  low cost |  not reliably sporicidal |  milk and dairy products  some heat-sensitive medical equipment |
| Tyndallization (Fractional Sterilization) |  heating to 80-100?C for 30 mins on successive days, with incubation periods in between |  resistant spores germinate and are killed on the second and third days |  time consuming  not reliably sporicidal |  heat sensitive materials such as bacteriologic media, solutions of chemicals, biological materials |
| Boiling |  maximum temperature obtainable is approximately 100?C 10-30 mins |  minimal equipment required |  cumbersome: not practical for everyday lab use  not reliably sporicidal |  small instruments and equipment |
| autoclaving |  steam under pressure  121?C/15 psi for 15-90 mins (gravity displacement autoclave)  132?C/27 psi for 4-20 minutes (pre-vacuum autoclave) |  minimal time required  most dependable sterilant for lab use |  loading and packing critical to performance  shielding dirt must first be removed  maintenance and quality control essential  damages heat-sensitive itmes |  penetration of sterile glassware, media and instruments  decontamination of reusable supplies and equipment  decontamination of infectious waste |

### 6.2.2 Other Physical Agents of Sterilization and Disinfection

#### 6.2.2.1 Ultraviolet Light (Germicidal Lamps)

The light (approximately 260 nm wavelength) emitted by UV lamps is germicidal, and can be used to reduce the number of pathogenic microorganisms on exposed surfaces and in air. However, UV light has poor penetrating power; accumulations of dust, dirt, grease or clumps of microorganisms may shield microorganisms from the direct exposure required for destruction. UV light presents skin and eye burn hazard, and factors such as lamp age and poor maintenance can reduce performance. For safe and reliable use of germicidal lamps:

- Clean the bulb at least every 2 weeks; turn off power and wipe with an alcohol-moistened cloth.

- Blue light output is not an indication of the lamp's effectiveness; measure radiation output at least twice yearly with a UV meter or replace the bulb when emission declines to 70% of its rated output.

- Post warning signs to discourage personnel from entering areas where there is potential exposure to UV light.

- Wear UV protective goggles, caps, gowns and gloves in rooms with UV installations.

#### 6.2.2.2 Miscellaneous Physical Methods

The procedures listed below are included for the reader's interest:

- Infrared radiation: used for heat treatment of small metal and glass items.

- Microwaves: used for treatment of liquids, nonmetallic objects, and biohazardous waste.

- Gamma irradiation: disrupts DNA and RNA in living organisms, and is used by hospital and laboratory suppliers for materials that do not tolerate heat and pressure (i.e., autoclaving) or chemical treatments.

- Membrane filtration: physically removes particulates (e.g., microorganisms) from heat-sensitive pharmaceutical and biological fluids. The size of the particles removed is determined by the pore size of the filter membrane.

## 6.3 Chemical Sterilants and Disinfectants

Instruments or materials which cannot withstand sterilization in a steam autoclave or dry-air oven can be sterilized with a gas such as ethylene oxide or a broad spectrum liquid chemical germicide. Chemical decontamination of surfaces may also be necessary for very large or fixed items. Since liquid chemical germicides generally require high concentrations and several hours of exposure time for sterilization purposes, they are usually used for disinfection rather than for sterilization purposes. The majority of chemical disinfectants have toxic properties: follow the manufacurer's directions for use and wear the appropriate personal protective equipment (e.g., gloves, eye protection, apron), especially when handling stock solutions.

Choice of a chemical germicide for use on contaminated equipment, supplies, laboratory surfaces or biohazardous waste depends upon a number of factors, including:

- number and nature of microbes to be destroyed (e.g., spores vs vegetative cells, bacteria vs viruses),

- type and configuration of item to be disinfected (fissures, crevices and enclosures may shield organisms),

- purpose of treatment (e.g., disinfection vs sterilization),

- interaction with other active chemicals,

- whether the item is covered with soil which might inactivate the disinfectant,

- contact time required for disinfection,

- toxicity to individuals, culture systems, environment, residual toxicity on items,

- pH, temperature, hardness of available dilution water,

- cost.

Direct contact between germicide and microorganism is essential for disinfection. Microorganisms can be shielded within air bubbles or under dirt, grease, oil, rust or clumps of microorganisms. Agar or proteinaceous nutrients and other cellular material can, either directly (through inactivation of the germicide) or indirectly (via physical shielding of microorganisms) reduce the efficacity of some liquid germicides.

No one chemical germicide is effective for all disinfection or sterilization purposes. A summary of chemical germicides, their use, effective concentrations, advantages and disadvantages can be found in Tables 3, 4A and 4B.

| TABLE 3 - Summary of concentrations used, contact times, advantages and disadvantages and uses of some of the halogen-releasing chemical germicides. The wide ranges of effective concentrations and contact times cited are due to a number of factors, including the interdependence of time and concentration, the variability in resistance of different microorganisms, the amount of organic material present and the desired effect (e.g., low-level vs high-level disinfection) |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Effective Concentrations, Contact Times | Advantages | Disadvantages | Examples of Uses |
| Chlorine Compounds: Sodium hypochlorite solution 1 (liquid bleach) |  100-10,000 ppm (.01-1%) free chlorine  10-60 minutes (>= 3,000 ppm for broad spectrum) |  broad spectrum  inexpensive  widely available  bactericidal at low temperature |  toxic, corrosive to skin and metals  unstable at optimum effective pH of 6  inactivated by organic matter  deteriorates under light and heat: shelf life of dilutions is less than 1 week |  general disinfectant  waste liquids  surface decontamination  emergency spill clean up  instrument disinfection |
| Calcium hypochlorite2 granules, powder, tablets |  as for liquid bleach |  as for liquid bleach but more stable |  as for liquid bleach above, except shelf life is longer |  as for liquid bleach |
| NaDCC3 (Sodium dichloroisocyanurate) powder, granules, tablets |  as for liquid bleach |  more stable than hypochlorites  stable at pH 6.0 |  toxic, corrosive  inactivated by organic matter |  as for liquid bleach |
| Chloramine-T4 (Sodium tosylchloramide) powder or tablets |  as for liquid bleach |  more stable, less affected by organic matter than hypochlorites  longer activity than hypochlorites |  deteriorates under humidity, light and heat |  as for liquid bleach |
| Chlorine dioxide5 |  demand-release of chlorine dioxide in situ |  longer activity than other chlorine compounds  less corrosive, less toxic than other chlorine compounds  effective at pH 6-10 |  aqueous solutions decompose under light |  instrument disinfection  gas sterilization of germ-free animal chambers |
| Iodine Preparations: Iodophors6 |  30-1,000 ppm (.003-.1%) free iodine  10-30 minutes |  broad spectrum  germicidal over a wide pH range  generally nonstaining, less toxic and less irritating than aqueous or alcoholic iodine solutions |  not consistently sporicidal  efficacy reduced by organic matter  some iodophor solutions support growth of Pseudomonas7 |  germicidal soaps and antiseptics  surface decontamination  work surface wipedown  instrument disinfection |

#### Notes:

1 a 1/10 dilution of 5.25% bleach provides 5,250 ppm available chlorine

2 "high tested" provides 70-72% available chlorine; chlorinated lime or bleaching powder provides approximately 35% available chlorine

3 appproximately 60% available chlorine

4 approximately 25% available chlorine

5 To avoid shipping of this extremely reactive product, reagents ("base" and "activator") from commercially available kits are mixed with water to generate chlorine dioxide immediately prior to use

6 10% povidone-iodine provides 1% available iodine

7 An iodophor stock solution may actually be a less effective germicide than its dilution. For example, a full-strength (10%) solution of povidone-iodine provides approximately 10 times less free available iodine than a 1/100 dilution. Iodophors must be used at the manufacturer's recommended concentrations.

| TABLE 4A Summary of recommended concentrations, contact times, advantages and disadvantages of non-halogen chemical germicides. The wide ranges of effective concentrations and contact times cited reflect the interdependence of time and concentration as well as factors such as resistance of the particular class or strain of target microorganism(s) and desired effect. Also, some germicides are available in combinations (e.g., glutaraldehyde/phenol or peracetic acid/alcohol mixtures) which are synergistic whereby the components in combination produce a greater antimicrobial effect than the sum of their individual effects. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Effective Concentrations and Contact Times | Advantages | Disadvantages | Examples of Laboratory Uses |
| Alcohols |  70-80% ethanol  60-95% isopropanol  10-30 minutes |  low toxicity  rapid action  low residue  non-corrosive |  rapid evaporation limits contact time  flammable, eye irratant  may damage rubber, plastic, shellac  ineffective against bacterial spores |  skin disinfectant (antiseptic)  surface decontamination  benchtop, cabinet wipedown |
| Phenolic Compounds |  400-50,000 ppm (.05-1.5%)  10-30 minutes |  tolerant of organic load, "hard" dilution water  leaves an active residue (may be desirable on some surfaces)  biodegradable |  pungent odour, corrosive, some forms toxic  not sporicidal; limited activity against viruses  leaves a residual film (undesirable in culture systems)  may support growth of bacteria1 |  instruments and equipment disinfection  disinfection of floors and other surfaces  antiseptic soaps and lotions |
| Quaternary Ammonium Compounds |  500-15,000 ppm (.05-1.5%)  10-30 minutes |  combined detergent and germicidal activity  stable  working dilutions have low toxicity |  non sporicidal, limited activity against viruses, mycobacteria  most formulations not readily biodegradable  may support growth of bacteria2 |  surface decontamination  equipment wipedown  antiseptic formulations available  floors and walls |
| Hydrogen Peroxide |  3-30% aqueous solution  10-60 minutes  6% for 30 minutes may kill spores |  rapid action  no residue  low toxicity  environmentally safe |  limited sporicidal activity  corrosive to some metals  potentially explosive at high concentrations  stock solutions irritating to skin and eyes |  surface decontamination  instruments and equipment |
| Peracetic Acid (PAA) |  .001-.3% aqueous solution  gas phase: 2-4%  5-120 minutes |  broad spectrum  sporicidal at low temperatures  can tolerate organic load  rapid action  nontoxic decomposition products  leaves no residue |  pungent odour  corrosive to some metals  shelf life of dilutions is less than 1 week  stock solutions irritating to skin and eyes  stock must be protected from heat, light  gas phase: respiratory irritant, fire hazard above 55?C |  instruments and equipment  gas phase sterilization of chambers for germ-free animals |

| Table 4B - Summary of recommended concentrations, contact times, advantages and disadvantages of non-halogen chemical germicides. The wide ranges of effective concentrations and contact times cited reflect the interdependence of time and concentration as well as factors such as resistance of the particular class or strain of target microorganism(s) and desired effect. Also, some germicides are available in combinations (e.g., glutaraldehyde/phenol or peracetic acid/alcohol mixtures) which are synergistic whereby the components in combination produce a greater antimicrobial effect than the sum of their individual effects. |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | Effective Concentrations and Contact Times | Advantages | Disadvantages | Examples of Laboratory Uses |
| Aldehydes: |  |  |  |  |
| Glutaraldehyde |  0.5-2.5% alkalinized aqueous solution  2-30 mins; up to 12 hours to kill all spores |  broad spectrum  does not corrode metal  can tolerate organic load |  expensive  pH, temperature dependent  pungent odour  toxic: skin, eye, respiratory tract irritant  activated solutions have less than 2-week shelf life |  cold sterilant and fixative  surface decontamination  instruments, equipment, glassware |
| Formalin (37% aqueous formaldehyde) |  3-27% formalin (1-10% formaldehyde) in 70-90% alcohol  10-30 minutes |  broad spectrum  inexpensive  does not corrode metal  can tolerate organic load |  pungent odour  skin, eye and respiratory tract irritant  potential carcinogen (animal studies)  may require 24 hrs or more to kill all spores |  cold sterilant and fixative  surface decontamination  instruments and equipment |
| Formaldehyde (gas) |  1-3 hours |  as for formalin  effective penetration |  as for formulin  flammable  porr penetration of covered surfaces |  on site decontamination of biological safety cabinet HEPA filters  enclosed areas |
| Ethylene Oxide Gas |  50-1200 mg/L  1-12 hours |  broad spectrum  no heat or moisture evolved  penetrates packaging materials |  flammable, reactive  toxic: potential carcinogen and mutagen  some sterilized items may need more than 24 hours for outgassing |  heat or moisture sensitive supplies, instruments and equipment |